Report for ctc510
In order to characterize LFR transfer function , this scenario generates 6 sinewaves without phases on one Analog discovery (3 signals on each output with T duplicators) for several freq (88 freqs @f0 + out of range freq) succesively and store results in different raw formatted files. ASM period is set to 4 sec and SWF to 16 sec. Each Acquisition is made in NORMAL mode during 26 seconds to ensure that we get at least 1 BP2 packet, several ASM packets and SWF packets. LFR enter NORMAL mode on a sharp second to ensure acquisition starts on a sharp second then a trigger is sent to Analog discovery at a user defined time to start generation of the signal.
Configuration

Parameter	Value
wave generator	analog discovery mapping:	B123_LF SN:210244516938 0	VE12_LF SN:210244516938 1
SocExplorer	0.6.2
LFRControlPlugin	3.0.0.1
VHDL	1.1.89
FSW	3.0.0.10
SP0 COMMON PARAM	0
SP1 COMMON PARAM	0
R0 COMMON PARAM	0
R1 COMMON PARAM	0
R2 COMMON PARAM	0

Scenario

Time	Step
16:08:22.740	This is /opt/CALIBRATION/CTC510/scenario
16:08:22.743	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:08:22.744	TC_LFR_LOAD_NORMAL_PAR *** set snapshot period to 16 seconds
16:08:22.746	TC_LFR_LOAD_NORMAL_PAR *** set asm period to 4 seconds
16:08:22.765	Configure sinewaves generation: 2.0Vpp @12.000 Hz
16:08:24.800	Time has been set to : 498150504
16:08:24.801	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498150505
16:08:24.803	Trigger will be sent to Analog Discovery at : 498150507
16:08:38.855	We start to log RAW and LOG files.
16:09:04.857	We stop to log RAW and LOG files.
16:09:06.859	2015_10_14_16_08_38_packet_record.data contains data at freq : 12.000Hz
16:09:06.861	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:09:08.875	Configure sinewaves generation: 2.0Vpp @24.000 Hz
16:09:10.925	Time has been set to : 498150550
16:09:10.927	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498150551
16:09:10.928	Trigger will be sent to Analog Discovery at : 498150553
16:09:24.970	We start to log RAW and LOG files.
16:09:50.973	We stop to log RAW and LOG files.
16:09:52.975	2015_10_14_16_09_24_packet_record.data contains data at freq : 24.000Hz
16:09:52.977	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:09:54.986	Configure sinewaves generation: 2.0Vpp @36.000 Hz
16:09:57.33	Time has been set to : 498150597
16:09:57.34	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498150598
16:09:57.36	Trigger will be sent to Analog Discovery at : 498150600
16:10:11.89	We start to log RAW and LOG files.
16:10:37.91	We stop to log RAW and LOG files.
16:10:39.92	2015_10_14_16_10_11_packet_record.data contains data at freq : 36.000Hz
16:10:39.94	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:10:41.103	Configure sinewaves generation: 2.0Vpp @48.000 Hz
16:10:43.146	Time has been set to : 498150643
16:10:43.148	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498150644
16:10:43.149	Trigger will be sent to Analog Discovery at : 498150646
16:10:57.191	We start to log RAW and LOG files.
16:11:23.193	We stop to log RAW and LOG files.
16:11:25.195	2015_10_14_16_10_57_packet_record.data contains data at freq : 48.000Hz
16:11:25.197	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:11:27.206	Configure sinewaves generation: 2.0Vpp @60.000 Hz
16:11:29.258	Time has been set to : 498150689
16:11:29.259	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498150690
16:11:29.261	Trigger will be sent to Analog Discovery at : 498150692
16:11:43.303	We start to log RAW and LOG files.
16:12:09.305	We stop to log RAW and LOG files.
16:12:11.307	2015_10_14_16_11_43_packet_record.data contains data at freq : 60.000Hz
16:12:11.308	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:12:13.317	Configure sinewaves generation: 2.0Vpp @72.000 Hz
16:12:15.359	Time has been set to : 498150735
16:12:15.360	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498150736
16:12:15.362	Trigger will be sent to Analog Discovery at : 498150738
16:12:29.406	We start to log RAW and LOG files.
16:12:55.408	We stop to log RAW and LOG files.
16:12:57.410	2015_10_14_16_12_29_packet_record.data contains data at freq : 72.000Hz
16:12:57.412	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:12:59.422	Configure sinewaves generation: 2.0Vpp @84.000 Hz
16:13:01.466	Time has been set to : 498150781
16:13:01.467	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498150782
16:13:01.469	Trigger will be sent to Analog Discovery at : 498150784
16:13:15.511	We start to log RAW and LOG files.
16:13:41.513	We stop to log RAW and LOG files.
16:13:43.515	2015_10_14_16_13_15_packet_record.data contains data at freq : 84.000Hz
16:13:43.517	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:13:45.528	Configure sinewaves generation: 2.0Vpp @96.000 Hz
16:13:47.578	Time has been set to : 498150827
16:13:47.580	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498150828
16:13:47.581	Trigger will be sent to Analog Discovery at : 498150830
16:14:01.623	We start to log RAW and LOG files.
16:14:27.625	We stop to log RAW and LOG files.
16:14:29.627	2015_10_14_16_14_01_packet_record.data contains data at freq : 96.000Hz
16:14:29.629	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:14:31.642	Configure sinewaves generation: 2.0Vpp @192.000 Hz
16:14:33.690	Time has been set to : 498150873
16:14:33.692	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498150874
16:14:33.693	Trigger will be sent to Analog Discovery at : 498150876
16:14:47.735	We start to log RAW and LOG files.
16:15:13.737	We stop to log RAW and LOG files.
16:15:15.739	2015_10_14_16_14_47_packet_record.data contains data at freq : 192.000Hz
16:15:15.741	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:15:17.753	Configure sinewaves generation: 2.0Vpp @288.000 Hz
16:15:19.804	Time has been set to : 498150919
16:15:19.806	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498150920
16:15:19.808	Trigger will be sent to Analog Discovery at : 498150922
16:15:33.849	We start to log RAW and LOG files.
16:15:59.851	We stop to log RAW and LOG files.
16:16:01.853	2015_10_14_16_15_33_packet_record.data contains data at freq : 288.000Hz
16:16:01.855	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:16:03.869	Configure sinewaves generation: 2.0Vpp @384.000 Hz
16:16:05.915	Time has been set to : 498150965
16:16:05.916	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498150966
16:16:05.918	Trigger will be sent to Analog Discovery at : 498150968
16:16:19.971	We start to log RAW and LOG files.
16:16:45.978	We stop to log RAW and LOG files.
16:16:47.980	2015_10_14_16_16_19_packet_record.data contains data at freq : 384.000Hz
16:16:47.982	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:16:49.991	Configure sinewaves generation: 2.0Vpp @480.000 Hz
16:16:52.33	Time has been set to : 498151012
16:16:52.35	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151013
16:16:52.36	Trigger will be sent to Analog Discovery at : 498151015
16:17:06.89	We start to log RAW and LOG files.
16:17:32.91	We stop to log RAW and LOG files.
16:17:34.93	2015_10_14_16_17_06_packet_record.data contains data at freq : 480.000Hz
16:17:34.95	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:17:36.104	Configure sinewaves generation: 2.0Vpp @576.000 Hz
16:17:38.146	Time has been set to : 498151058
16:17:38.148	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151059
16:17:38.150	Trigger will be sent to Analog Discovery at : 498151061
16:17:52.203	We start to log RAW and LOG files.
16:18:18.206	We stop to log RAW and LOG files.
16:18:20.208	2015_10_14_16_17_52_packet_record.data contains data at freq : 576.000Hz
16:18:20.210	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:18:22.217	Configure sinewaves generation: 2.0Vpp @672.000 Hz
16:18:24.269	Time has been set to : 498151104
16:18:24.271	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151105
16:18:24.272	Trigger will be sent to Analog Discovery at : 498151107
16:18:38.325	We start to log RAW and LOG files.
16:19:04.327	We stop to log RAW and LOG files.
16:19:06.329	2015_10_14_16_18_38_packet_record.data contains data at freq : 672.000Hz
16:19:06.331	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:19:08.339	Configure sinewaves generation: 2.0Vpp @768.000 Hz
16:19:10.388	Time has been set to : 498151150
16:19:10.389	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151151
16:19:10.391	Trigger will be sent to Analog Discovery at : 498151153
16:19:24.444	We start to log RAW and LOG files.
16:19:50.446	We stop to log RAW and LOG files.
16:19:52.448	2015_10_14_16_19_24_packet_record.data contains data at freq : 768.000Hz
16:19:52.450	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:19:54.458	Configure sinewaves generation: 2.0Vpp @864.000 Hz
16:19:56.499	Time has been set to : 498151196
16:19:56.501	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151197
16:19:56.502	Trigger will be sent to Analog Discovery at : 498151199
16:20:10.555	We start to log RAW and LOG files.
16:20:36.557	We stop to log RAW and LOG files.
16:20:38.559	2015_10_14_16_20_10_packet_record.data contains data at freq : 864.000Hz
16:20:38.561	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:20:40.578	Configure sinewaves generation: 2.0Vpp @960.000 Hz
16:20:42.625	Time has been set to : 498151242
16:20:42.627	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151243
16:20:42.629	Trigger will be sent to Analog Discovery at : 498151245
16:20:56.681	We start to log RAW and LOG files.
16:21:22.683	We stop to log RAW and LOG files.
16:21:24.685	2015_10_14_16_20_56_packet_record.data contains data at freq : 960.000Hz
16:21:24.687	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:21:26.701	Configure sinewaves generation: 2.0Vpp @1056.000 Hz
16:21:28.758	Time has been set to : 498151288
16:21:28.760	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151289
16:21:28.762	Trigger will be sent to Analog Discovery at : 498151291
16:21:42.803	We start to log RAW and LOG files.
16:22:08.805	We stop to log RAW and LOG files.
16:22:10.807	2015_10_14_16_21_42_packet_record.data contains data at freq : 1056.000Hz
16:22:10.809	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:22:12.826	Configure sinewaves generation: 2.0Vpp @1152.000 Hz
16:22:14.867	Time has been set to : 498151334
16:22:14.869	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151335
16:22:14.870	Trigger will be sent to Analog Discovery at : 498151337
16:22:28.912	We start to log RAW and LOG files.
16:22:54.920	We stop to log RAW and LOG files.
16:22:56.922	2015_10_14_16_22_28_packet_record.data contains data at freq : 1152.000Hz
16:22:56.924	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:22:58.936	Configure sinewaves generation: 2.0Vpp @1248.000 Hz
16:23:00.984	Time has been set to : 498151380
16:23:00.985	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151381
16:23:00.987	Trigger will be sent to Analog Discovery at : 498151383
16:23:15.33	We start to log RAW and LOG files.
16:23:41.35	We stop to log RAW and LOG files.
16:23:43.37	2015_10_14_16_23_15_packet_record.data contains data at freq : 1248.000Hz
16:23:43.39	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:23:45.46	Configure sinewaves generation: 2.0Vpp @1344.000 Hz
16:23:47.99	Time has been set to : 498151427
16:23:47.101	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151428
16:23:47.103	Trigger will be sent to Analog Discovery at : 498151430
16:24:01.144	We start to log RAW and LOG files.
16:24:27.146	We stop to log RAW and LOG files.
16:24:29.148	2015_10_14_16_24_01_packet_record.data contains data at freq : 1344.000Hz
16:24:29.150	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:24:31.159	Configure sinewaves generation: 2.0Vpp @1440.000 Hz
16:24:33.213	Time has been set to : 498151473
16:24:33.215	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151474
16:24:33.217	Trigger will be sent to Analog Discovery at : 498151476
16:24:47.258	We start to log RAW and LOG files.
16:25:13.260	We stop to log RAW and LOG files.
16:25:15.262	2015_10_14_16_24_47_packet_record.data contains data at freq : 1440.000Hz
16:25:15.264	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:25:17.273	Configure sinewaves generation: 2.0Vpp @1536.000 Hz
16:25:19.326	Time has been set to : 498151519
16:25:19.327	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151520
16:25:19.329	Trigger will be sent to Analog Discovery at : 498151522
16:25:33.377	We start to log RAW and LOG files.
16:25:59.379	We stop to log RAW and LOG files.
16:26:01.381	2015_10_14_16_25_33_packet_record.data contains data at freq : 1536.000Hz
16:26:01.383	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:26:03.391	Configure sinewaves generation: 2.0Vpp @1632.000 Hz
16:26:05.444	Time has been set to : 498151565
16:26:05.446	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151566
16:26:05.448	Trigger will be sent to Analog Discovery at : 498151568
16:26:19.501	We start to log RAW and LOG files.
16:26:45.503	We stop to log RAW and LOG files.
16:26:47.505	2015_10_14_16_26_19_packet_record.data contains data at freq : 1632.000Hz
16:26:47.507	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:26:49.515	Configure sinewaves generation: 2.0Vpp @1728.000 Hz
16:26:51.557	Time has been set to : 498151611
16:26:51.559	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151612
16:26:51.560	Trigger will be sent to Analog Discovery at : 498151614
16:27:05.613	We start to log RAW and LOG files.
16:27:31.615	We stop to log RAW and LOG files.
16:27:33.617	2015_10_14_16_27_05_packet_record.data contains data at freq : 1728.000Hz
16:27:33.619	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:27:35.626	Configure sinewaves generation: 2.0Vpp @1824.000 Hz
16:27:37.669	Time has been set to : 498151657
16:27:37.671	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151658
16:27:37.672	Trigger will be sent to Analog Discovery at : 498151660
16:27:51.725	We start to log RAW and LOG files.
16:28:17.727	We stop to log RAW and LOG files.
16:28:19.729	2015_10_14_16_27_51_packet_record.data contains data at freq : 1824.000Hz
16:28:19.731	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:28:21.739	Configure sinewaves generation: 2.0Vpp @1920.000 Hz
16:28:23.793	Time has been set to : 498151703
16:28:23.794	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151704
16:28:23.796	Trigger will be sent to Analog Discovery at : 498151706
16:28:37.849	We start to log RAW and LOG files.
16:29:03.851	We stop to log RAW and LOG files.
16:29:05.853	2015_10_14_16_28_37_packet_record.data contains data at freq : 1920.000Hz
16:29:05.855	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:29:07.862	Configure sinewaves generation: 2.0Vpp @2016.000 Hz
16:29:09.917	Time has been set to : 498151749
16:29:09.919	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151750
16:29:09.920	Trigger will be sent to Analog Discovery at : 498151752
16:29:23.962	We start to log RAW and LOG files.
16:29:49.967	We stop to log RAW and LOG files.
16:29:51.970	2015_10_14_16_29_23_packet_record.data contains data at freq : 2016.000Hz
16:29:51.972	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:29:53.985	Configure sinewaves generation: 2.0Vpp @2112.000 Hz
16:29:56.34	Time has been set to : 498151796
16:29:56.35	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151797
16:29:56.37	Trigger will be sent to Analog Discovery at : 498151799
16:30:10.90	We start to log RAW and LOG files.
16:30:36.92	We stop to log RAW and LOG files.
16:30:38.94	2015_10_14_16_30_10_packet_record.data contains data at freq : 2112.000Hz
16:30:38.96	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:30:40.102	Configure sinewaves generation: 2.0Vpp @2208.000 Hz
16:30:42.145	Time has been set to : 498151842
16:30:42.146	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151843
16:30:42.148	Trigger will be sent to Analog Discovery at : 498151845
16:30:56.190	We start to log RAW and LOG files.
16:31:22.192	We stop to log RAW and LOG files.
16:31:24.194	2015_10_14_16_30_56_packet_record.data contains data at freq : 2208.000Hz
16:31:24.196	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:31:26.207	Configure sinewaves generation: 2.0Vpp @2304.000 Hz
16:31:28.258	Time has been set to : 498151888
16:31:28.259	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151889
16:31:28.261	Trigger will be sent to Analog Discovery at : 498151891
16:31:42.314	We start to log RAW and LOG files.
16:32:08.316	We stop to log RAW and LOG files.
16:32:10.318	2015_10_14_16_31_42_packet_record.data contains data at freq : 2304.000Hz
16:32:10.320	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:32:12.327	Configure sinewaves generation: 2.0Vpp @2400.000 Hz
16:32:14.366	Time has been set to : 498151934
16:32:14.368	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151935
16:32:14.370	Trigger will be sent to Analog Discovery at : 498151937
16:32:28.408	We start to log RAW and LOG files.
16:32:54.410	We stop to log RAW and LOG files.
16:32:56.412	2015_10_14_16_32_28_packet_record.data contains data at freq : 2400.000Hz
16:32:56.414	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:32:58.426	Configure sinewaves generation: 2.0Vpp @2496.000 Hz
16:33:00.480	Time has been set to : 498151980
16:33:00.481	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498151981
16:33:00.483	Trigger will be sent to Analog Discovery at : 498151983
16:33:14.525	We start to log RAW and LOG files.
16:33:40.527	We stop to log RAW and LOG files.
16:33:42.529	2015_10_14_16_33_14_packet_record.data contains data at freq : 2496.000Hz
16:33:42.531	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:33:44.538	Configure sinewaves generation: 2.0Vpp @2592.000 Hz
16:33:46.579	Time has been set to : 498152026
16:33:46.581	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152027
16:33:46.583	Trigger will be sent to Analog Discovery at : 498152029
16:34:00.636	We start to log RAW and LOG files.
16:34:26.638	We stop to log RAW and LOG files.
16:34:28.640	2015_10_14_16_34_00_packet_record.data contains data at freq : 2592.000Hz
16:34:28.642	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:34:30.653	Configure sinewaves generation: 2.0Vpp @2688.000 Hz
16:34:32.706	Time has been set to : 498152072
16:34:32.708	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152073
16:34:32.710	Trigger will be sent to Analog Discovery at : 498152075
16:34:46.752	We start to log RAW and LOG files.
16:35:12.754	We stop to log RAW and LOG files.
16:35:14.756	2015_10_14_16_34_46_packet_record.data contains data at freq : 2688.000Hz
16:35:14.758	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:35:16.775	Configure sinewaves generation: 2.0Vpp @2784.000 Hz
16:35:18.826	Time has been set to : 498152118
16:35:18.828	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152119
16:35:18.830	Trigger will be sent to Analog Discovery at : 498152121
16:35:32.882	We start to log RAW and LOG files.
16:35:58.884	We stop to log RAW and LOG files.
16:36:00.887	2015_10_14_16_35_32_packet_record.data contains data at freq : 2784.000Hz
16:36:00.889	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:36:02.899	Configure sinewaves generation: 2.0Vpp @2880.000 Hz
16:36:04.951	Time has been set to : 498152164
16:36:04.953	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152165
16:36:04.955	Trigger will be sent to Analog Discovery at : 498152167
16:36:18.996	We start to log RAW and LOG files.
16:36:45.0	We stop to log RAW and LOG files.
16:36:47.3	2015_10_14_16_36_18_packet_record.data contains data at freq : 2880.000Hz
16:36:47.5	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:36:49.13	Configure sinewaves generation: 2.0Vpp @2976.000 Hz
16:36:51.56	Time has been set to : 498152211
16:36:51.57	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152212
16:36:51.59	Trigger will be sent to Analog Discovery at : 498152214
16:37:05.112	We start to log RAW and LOG files.
16:37:31.114	We stop to log RAW and LOG files.
16:37:33.116	2015_10_14_16_37_05_packet_record.data contains data at freq : 2976.000Hz
16:37:33.119	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:37:35.125	Configure sinewaves generation: 2.0Vpp @3072.000 Hz
16:37:37.168	Time has been set to : 498152257
16:37:37.170	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152258
16:37:37.172	Trigger will be sent to Analog Discovery at : 498152260
16:37:51.224	We start to log RAW and LOG files.
16:38:17.227	We stop to log RAW and LOG files.
16:38:19.229	2015_10_14_16_37_51_packet_record.data contains data at freq : 3072.000Hz
16:38:19.231	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:38:21.238	Configure sinewaves generation: 2.0Vpp @3168.000 Hz
16:38:23.291	Time has been set to : 498152303
16:38:23.293	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152304
16:38:23.295	Trigger will be sent to Analog Discovery at : 498152306
16:38:37.336	We start to log RAW and LOG files.
16:39:03.338	We stop to log RAW and LOG files.
16:39:05.341	2015_10_14_16_38_37_packet_record.data contains data at freq : 3168.000Hz
16:39:05.343	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:39:07.351	Configure sinewaves generation: 2.0Vpp @3264.000 Hz
16:39:09.401	Time has been set to : 498152349
16:39:09.403	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152350
16:39:09.405	Trigger will be sent to Analog Discovery at : 498152352
16:39:23.446	We start to log RAW and LOG files.
16:39:49.448	We stop to log RAW and LOG files.
16:39:51.451	2015_10_14_16_39_23_packet_record.data contains data at freq : 3264.000Hz
16:39:51.453	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:39:53.460	Configure sinewaves generation: 2.0Vpp @3360.000 Hz
16:39:55.514	Time has been set to : 498152395
16:39:55.515	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152396
16:39:55.517	Trigger will be sent to Analog Discovery at : 498152398
16:40:09.559	We start to log RAW and LOG files.
16:40:35.561	We stop to log RAW and LOG files.
16:40:37.563	2015_10_14_16_40_09_packet_record.data contains data at freq : 3360.000Hz
16:40:37.565	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:40:39.576	Configure sinewaves generation: 2.0Vpp @3456.000 Hz
16:40:41.626	Time has been set to : 498152441
16:40:41.628	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152442
16:40:41.630	Trigger will be sent to Analog Discovery at : 498152444
16:40:55.671	We start to log RAW and LOG files.
16:41:21.673	We stop to log RAW and LOG files.
16:41:23.675	2015_10_14_16_40_55_packet_record.data contains data at freq : 3456.000Hz
16:41:23.677	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:41:25.688	Configure sinewaves generation: 2.0Vpp @3552.000 Hz
16:41:27.738	Time has been set to : 498152487
16:41:27.740	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152488
16:41:27.741	Trigger will be sent to Analog Discovery at : 498152490
16:41:41.783	We start to log RAW and LOG files.
16:42:07.785	We stop to log RAW and LOG files.
16:42:09.788	2015_10_14_16_41_41_packet_record.data contains data at freq : 3552.000Hz
16:42:09.790	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:42:11.800	Configure sinewaves generation: 2.0Vpp @3648.000 Hz
16:42:13.850	Time has been set to : 498152533
16:42:13.852	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152534
16:42:13.854	Trigger will be sent to Analog Discovery at : 498152536
16:42:27.895	We start to log RAW and LOG files.
16:42:53.898	We stop to log RAW and LOG files.
16:42:55.902	2015_10_14_16_42_27_packet_record.data contains data at freq : 3648.000Hz
16:42:55.904	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:42:57.915	Configure sinewaves generation: 2.0Vpp @3744.000 Hz
16:42:59.974	Time has been set to : 498152579
16:42:59.976	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152580
16:42:59.978	Trigger will be sent to Analog Discovery at : 498152582
16:43:14.22	We start to log RAW and LOG files.
16:43:40.25	We stop to log RAW and LOG files.
16:43:42.28	2015_10_14_16_43_14_packet_record.data contains data at freq : 3744.000Hz
16:43:42.30	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:43:44.38	Configure sinewaves generation: 2.0Vpp @3840.000 Hz
16:43:46.90	Time has been set to : 498152626
16:43:46.91	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152627
16:43:46.93	Trigger will be sent to Analog Discovery at : 498152629
16:44:00.146	We start to log RAW and LOG files.
16:44:26.148	We stop to log RAW and LOG files.
16:44:28.151	2015_10_14_16_44_00_packet_record.data contains data at freq : 3840.000Hz
16:44:28.153	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:44:30.161	Configure sinewaves generation: 2.0Vpp @3936.000 Hz
16:44:32.213	Time has been set to : 498152672
16:44:32.215	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152673
16:44:32.216	Trigger will be sent to Analog Discovery at : 498152675
16:44:46.269	We start to log RAW and LOG files.
16:45:12.271	We stop to log RAW and LOG files.
16:45:14.274	2015_10_14_16_44_46_packet_record.data contains data at freq : 3936.000Hz
16:45:14.276	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:45:16.284	Configure sinewaves generation: 2.0Vpp @4032.000 Hz
16:45:18.336	Time has been set to : 498152718
16:45:18.338	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152719
16:45:18.340	Trigger will be sent to Analog Discovery at : 498152721
16:45:32.399	We start to log RAW and LOG files.
16:45:58.401	We stop to log RAW and LOG files.
16:46:00.404	2015_10_14_16_45_32_packet_record.data contains data at freq : 4032.000Hz
16:46:00.406	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:46:02.416	Configure sinewaves generation: 2.0Vpp @4128.000 Hz
16:46:04.458	Time has been set to : 498152764
16:46:04.460	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152765
16:46:04.461	Trigger will be sent to Analog Discovery at : 498152767
16:46:18.514	We start to log RAW and LOG files.
16:46:44.516	We stop to log RAW and LOG files.
16:46:46.519	2015_10_14_16_46_18_packet_record.data contains data at freq : 4128.000Hz
16:46:46.521	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:46:48.529	Configure sinewaves generation: 2.0Vpp @4224.000 Hz
16:46:50.570	Time has been set to : 498152810
16:46:50.572	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152811
16:46:50.574	Trigger will be sent to Analog Discovery at : 498152813
16:47:04.626	We start to log RAW and LOG files.
16:47:30.629	We stop to log RAW and LOG files.
16:47:32.631	2015_10_14_16_47_04_packet_record.data contains data at freq : 4224.000Hz
16:47:32.633	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:47:34.646	Configure sinewaves generation: 2.0Vpp @4320.000 Hz
16:47:36.693	Time has been set to : 498152856
16:47:36.695	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152857
16:47:36.697	Trigger will be sent to Analog Discovery at : 498152859
16:47:50.749	We start to log RAW and LOG files.
16:48:16.751	We stop to log RAW and LOG files.
16:48:18.754	2015_10_14_16_47_50_packet_record.data contains data at freq : 4320.000Hz
16:48:18.756	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:48:20.770	Configure sinewaves generation: 2.0Vpp @4416.000 Hz
16:48:22.818	Time has been set to : 498152902
16:48:22.820	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152903
16:48:22.822	Trigger will be sent to Analog Discovery at : 498152905
16:48:36.874	We start to log RAW and LOG files.
16:49:02.876	We stop to log RAW and LOG files.
16:49:04.879	2015_10_14_16_48_36_packet_record.data contains data at freq : 4416.000Hz
16:49:04.881	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:49:06.894	Configure sinewaves generation: 2.0Vpp @4512.000 Hz
16:49:08.946	Time has been set to : 498152948
16:49:08.948	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152949
16:49:08.950	Trigger will be sent to Analog Discovery at : 498152951
16:49:22.993	We start to log RAW and LOG files.
16:49:48.999	We stop to log RAW and LOG files.
16:49:51.2	2015_10_14_16_49_22_packet_record.data contains data at freq : 4512.000Hz
16:49:51.4	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:49:53.14	Configure sinewaves generation: 2.0Vpp @4608.000 Hz
16:49:55.67	Time has been set to : 498152995
16:49:55.68	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498152996
16:49:55.70	Trigger will be sent to Analog Discovery at : 498152998
16:50:09.112	We start to log RAW and LOG files.
16:50:35.114	We stop to log RAW and LOG files.
16:50:37.116	2015_10_14_16_50_09_packet_record.data contains data at freq : 4608.000Hz
16:50:37.118	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:50:39.125	Configure sinewaves generation: 2.0Vpp @4704.000 Hz
16:50:41.169	Time has been set to : 498153041
16:50:41.171	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153042
16:50:41.172	Trigger will be sent to Analog Discovery at : 498153044
16:50:55.214	We start to log RAW and LOG files.
16:51:21.216	We stop to log RAW and LOG files.
16:51:23.218	2015_10_14_16_50_55_packet_record.data contains data at freq : 4704.000Hz
16:51:23.220	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:51:25.228	Configure sinewaves generation: 2.0Vpp @4800.000 Hz
16:51:27.280	Time has been set to : 498153087
16:51:27.282	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153088
16:51:27.284	Trigger will be sent to Analog Discovery at : 498153090
16:51:41.325	We start to log RAW and LOG files.
16:52:07.327	We stop to log RAW and LOG files.
16:52:09.330	2015_10_14_16_51_41_packet_record.data contains data at freq : 4800.000Hz
16:52:09.332	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:52:11.342	Configure sinewaves generation: 2.0Vpp @4896.000 Hz
16:52:13.391	Time has been set to : 498153133
16:52:13.393	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153134
16:52:13.395	Trigger will be sent to Analog Discovery at : 498153136
16:52:27.447	We start to log RAW and LOG files.
16:52:53.449	We stop to log RAW and LOG files.
16:52:55.453	2015_10_14_16_52_27_packet_record.data contains data at freq : 4896.000Hz
16:52:55.455	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:52:57.462	Configure sinewaves generation: 2.0Vpp @4992.000 Hz
16:52:59.515	Time has been set to : 498153179
16:52:59.517	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153180
16:52:59.519	Trigger will be sent to Analog Discovery at : 498153182
16:53:13.560	We start to log RAW and LOG files.
16:53:39.562	We stop to log RAW and LOG files.
16:53:41.565	2015_10_14_16_53_13_packet_record.data contains data at freq : 4992.000Hz
16:53:41.567	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:53:43.582	Configure sinewaves generation: 2.0Vpp @5088.000 Hz
16:53:45.641	Time has been set to : 498153225
16:53:45.643	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153226
16:53:45.645	Trigger will be sent to Analog Discovery at : 498153228
16:53:59.686	We start to log RAW and LOG files.
16:54:25.688	We stop to log RAW and LOG files.
16:54:27.691	2015_10_14_16_53_59_packet_record.data contains data at freq : 5088.000Hz
16:54:27.693	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:54:29.706	Configure sinewaves generation: 2.0Vpp @5184.000 Hz
16:54:31.754	Time has been set to : 498153271
16:54:31.756	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153272
16:54:31.758	Trigger will be sent to Analog Discovery at : 498153274
16:54:45.811	We start to log RAW and LOG files.
16:55:11.813	We stop to log RAW and LOG files.
16:55:13.815	2015_10_14_16_54_45_packet_record.data contains data at freq : 5184.000Hz
16:55:13.817	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:55:15.829	Configure sinewaves generation: 2.0Vpp @5280.000 Hz
16:55:17.862	Time has been set to : 498153317
16:55:17.864	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153318
16:55:17.866	Trigger will be sent to Analog Discovery at : 498153320
16:55:31.919	We start to log RAW and LOG files.
16:55:57.927	We stop to log RAW and LOG files.
16:55:59.930	2015_10_14_16_55_31_packet_record.data contains data at freq : 5280.000Hz
16:55:59.932	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:56:01.940	Configure sinewaves generation: 2.0Vpp @5376.000 Hz
16:56:03.987	Time has been set to : 498153363
16:56:03.989	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153364
16:56:03.991	Trigger will be sent to Analog Discovery at : 498153366
16:56:18.45	We start to log RAW and LOG files.
16:56:44.47	We stop to log RAW and LOG files.
16:56:46.50	2015_10_14_16_56_18_packet_record.data contains data at freq : 5376.000Hz
16:56:46.52	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:56:48.60	Configure sinewaves generation: 2.0Vpp @5472.000 Hz
16:56:50.111	Time has been set to : 498153410
16:56:50.113	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153411
16:56:50.115	Trigger will be sent to Analog Discovery at : 498153413
16:57:04.167	We start to log RAW and LOG files.
16:57:30.169	We stop to log RAW and LOG files.
16:57:32.172	2015_10_14_16_57_04_packet_record.data contains data at freq : 5472.000Hz
16:57:32.174	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:57:34.182	Configure sinewaves generation: 2.0Vpp @5568.000 Hz
16:57:36.223	Time has been set to : 498153456
16:57:36.225	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153457
16:57:36.227	Trigger will be sent to Analog Discovery at : 498153459
16:57:50.268	We start to log RAW and LOG files.
16:58:16.271	We stop to log RAW and LOG files.
16:58:18.274	2015_10_14_16_57_50_packet_record.data contains data at freq : 5568.000Hz
16:58:18.276	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:58:20.293	Configure sinewaves generation: 2.0Vpp @5664.000 Hz
16:58:22.337	Time has been set to : 498153502
16:58:22.339	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153503
16:58:22.341	Trigger will be sent to Analog Discovery at : 498153505
16:58:36.400	We start to log RAW and LOG files.
16:59:02.403	We stop to log RAW and LOG files.
16:59:04.405	2015_10_14_16_58_36_packet_record.data contains data at freq : 5664.000Hz
16:59:04.407	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:59:06.417	Configure sinewaves generation: 2.0Vpp @5760.000 Hz
16:59:08.471	Time has been set to : 498153548
16:59:08.473	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153549
16:59:08.475	Trigger will be sent to Analog Discovery at : 498153551
16:59:22.516	We start to log RAW and LOG files.
16:59:48.518	We stop to log RAW and LOG files.
16:59:50.521	2015_10_14_16_59_22_packet_record.data contains data at freq : 5760.000Hz
16:59:50.523	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
16:59:52.531	Configure sinewaves generation: 2.0Vpp @5856.000 Hz
16:59:54.572	Time has been set to : 498153594
16:59:54.574	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153595
16:59:54.575	Trigger will be sent to Analog Discovery at : 498153597
17:00:08.628	We start to log RAW and LOG files.
17:00:34.630	We stop to log RAW and LOG files.
17:00:36.633	2015_10_14_17_00_08_packet_record.data contains data at freq : 5856.000Hz
17:00:36.635	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:00:38.652	Configure sinewaves generation: 2.0Vpp @5952.000 Hz
17:00:40.710	Time has been set to : 498153640
17:00:40.711	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153641
17:00:40.713	Trigger will be sent to Analog Discovery at : 498153643
17:00:54.755	We start to log RAW and LOG files.
17:01:20.757	We stop to log RAW and LOG files.
17:01:22.760	2015_10_14_17_00_54_packet_record.data contains data at freq : 5952.000Hz
17:01:22.762	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:01:24.779	Configure sinewaves generation: 2.0Vpp @6048.000 Hz
17:01:26.831	Time has been set to : 498153686
17:01:26.833	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153687
17:01:26.835	Trigger will be sent to Analog Discovery at : 498153689
17:01:40.876	We start to log RAW and LOG files.
17:02:06.878	We stop to log RAW and LOG files.
17:02:08.881	2015_10_14_17_01_40_packet_record.data contains data at freq : 6048.000Hz
17:02:08.883	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:02:10.891	Configure sinewaves generation: 2.0Vpp @6144.000 Hz
17:02:12.935	Time has been set to : 498153732
17:02:12.937	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153733
17:02:12.939	Trigger will be sent to Analog Discovery at : 498153735
17:02:26.991	We start to log RAW and LOG files.
17:02:52.995	We stop to log RAW and LOG files.
17:02:54.998	2015_10_14_17_02_26_packet_record.data contains data at freq : 6144.000Hz
17:02:55.0	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:02:57.13	Configure sinewaves generation: 2.0Vpp @6240.000 Hz
17:02:59.67	Time has been set to : 498153779
17:02:59.69	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153780
17:02:59.71	Trigger will be sent to Analog Discovery at : 498153782
17:03:13.112	We start to log RAW and LOG files.
17:03:39.114	We stop to log RAW and LOG files.
17:03:41.117	2015_10_14_17_03_13_packet_record.data contains data at freq : 6240.000Hz
17:03:41.119	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:03:43.126	Configure sinewaves generation: 2.0Vpp @6336.000 Hz
17:03:45.179	Time has been set to : 498153825
17:03:45.181	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153826
17:03:45.183	Trigger will be sent to Analog Discovery at : 498153828
17:03:59.224	We start to log RAW and LOG files.
17:04:25.227	We stop to log RAW and LOG files.
17:04:27.229	2015_10_14_17_03_59_packet_record.data contains data at freq : 6336.000Hz
17:04:27.232	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:04:29.240	Configure sinewaves generation: 2.0Vpp @6432.000 Hz
17:04:31.281	Time has been set to : 498153871
17:04:31.283	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153872
17:04:31.285	Trigger will be sent to Analog Discovery at : 498153874
17:04:45.338	We start to log RAW and LOG files.
17:05:11.340	We stop to log RAW and LOG files.
17:05:13.343	2015_10_14_17_04_45_packet_record.data contains data at freq : 6432.000Hz
17:05:13.345	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:05:15.362	Configure sinewaves generation: 2.0Vpp @6528.000 Hz
17:05:17.404	Time has been set to : 498153917
17:05:17.406	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153918
17:05:17.408	Trigger will be sent to Analog Discovery at : 498153920
17:05:31.460	We start to log RAW and LOG files.
17:05:57.462	We stop to log RAW and LOG files.
17:05:59.465	2015_10_14_17_05_31_packet_record.data contains data at freq : 6528.000Hz
17:05:59.467	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:06:01.475	Configure sinewaves generation: 2.0Vpp @6624.000 Hz
17:06:03.528	Time has been set to : 498153963
17:06:03.530	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498153964
17:06:03.532	Trigger will be sent to Analog Discovery at : 498153966
17:06:17.573	We start to log RAW and LOG files.
17:06:43.576	We stop to log RAW and LOG files.
17:06:45.578	2015_10_14_17_06_17_packet_record.data contains data at freq : 6624.000Hz
17:06:45.581	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:06:47.598	Configure sinewaves generation: 2.0Vpp @6720.000 Hz
17:06:49.640	Time has been set to : 498154009
17:06:49.642	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154010
17:06:49.644	Trigger will be sent to Analog Discovery at : 498154012
17:07:03.685	We start to log RAW and LOG files.
17:07:29.687	We stop to log RAW and LOG files.
17:07:31.690	2015_10_14_17_07_03_packet_record.data contains data at freq : 6720.000Hz
17:07:31.692	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:07:33.701	Configure sinewaves generation: 2.0Vpp @6816.000 Hz
17:07:35.752	Time has been set to : 498154055
17:07:35.754	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154056
17:07:35.756	Trigger will be sent to Analog Discovery at : 498154058
17:07:49.798	We start to log RAW and LOG files.
17:08:15.800	We stop to log RAW and LOG files.
17:08:17.803	2015_10_14_17_07_49_packet_record.data contains data at freq : 6816.000Hz
17:08:17.805	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:08:19.812	Configure sinewaves generation: 2.0Vpp @6912.000 Hz
17:08:21.853	Time has been set to : 498154101
17:08:21.855	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154102
17:08:21.857	Trigger will be sent to Analog Discovery at : 498154104
17:08:35.909	We start to log RAW and LOG files.
17:09:01.912	We stop to log RAW and LOG files.
17:09:03.915	2015_10_14_17_08_35_packet_record.data contains data at freq : 6912.000Hz
17:09:03.917	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:09:05.925	Configure sinewaves generation: 2.0Vpp @7008.000 Hz
17:09:07.976	Time has been set to : 498154147
17:09:07.978	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154148
17:09:07.979	Trigger will be sent to Analog Discovery at : 498154150
17:09:22.22	We start to log RAW and LOG files.
17:09:48.28	We stop to log RAW and LOG files.
17:09:50.30	2015_10_14_17_09_22_packet_record.data contains data at freq : 7008.000Hz
17:09:50.33	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:09:52.47	Configure sinewaves generation: 2.0Vpp @7104.000 Hz
17:09:54.90	Time has been set to : 498154194
17:09:54.91	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154195
17:09:54.93	Trigger will be sent to Analog Discovery at : 498154197
17:10:08.146	We start to log RAW and LOG files.
17:10:34.148	We stop to log RAW and LOG files.
17:10:36.151	2015_10_14_17_10_08_packet_record.data contains data at freq : 7104.000Hz
17:10:36.153	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:10:38.171	Configure sinewaves generation: 2.0Vpp @7200.000 Hz
17:10:40.213	Time has been set to : 498154240
17:10:40.215	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154241
17:10:40.217	Trigger will be sent to Analog Discovery at : 498154243
17:10:54.269	We start to log RAW and LOG files.
17:11:20.271	We stop to log RAW and LOG files.
17:11:22.274	2015_10_14_17_10_54_packet_record.data contains data at freq : 7200.000Hz
17:11:22.277	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:11:24.295	Configure sinewaves generation: 2.0Vpp @7296.000 Hz
17:11:26.347	Time has been set to : 498154286
17:11:26.349	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154287
17:11:26.351	Trigger will be sent to Analog Discovery at : 498154289
17:11:40.402	We start to log RAW and LOG files.
17:12:06.404	We stop to log RAW and LOG files.
17:12:08.407	2015_10_14_17_11_40_packet_record.data contains data at freq : 7296.000Hz
17:12:08.409	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:12:10.419	Configure sinewaves generation: 2.0Vpp @7392.000 Hz
17:12:12.462	Time has been set to : 498154332
17:12:12.464	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154333
17:12:12.466	Trigger will be sent to Analog Discovery at : 498154335
17:12:26.518	We start to log RAW and LOG files.
17:12:52.520	We stop to log RAW and LOG files.
17:12:54.523	2015_10_14_17_12_26_packet_record.data contains data at freq : 7392.000Hz
17:12:54.525	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:12:56.544	Configure sinewaves generation: 2.0Vpp @7488.000 Hz
17:12:58.596	Time has been set to : 498154378
17:12:58.598	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154379
17:12:58.600	Trigger will be sent to Analog Discovery at : 498154381
17:13:12.641	We start to log RAW and LOG files.
17:13:38.643	We stop to log RAW and LOG files.
17:13:40.646	2015_10_14_17_13_12_packet_record.data contains data at freq : 7488.000Hz
17:13:40.649	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:13:42.660	Configure sinewaves generation: 2.0Vpp @7584.000 Hz
17:13:44.708	Time has been set to : 498154424
17:13:44.710	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154425
17:13:44.712	Trigger will be sent to Analog Discovery at : 498154427
17:13:58.764	We start to log RAW and LOG files.
17:14:24.767	We stop to log RAW and LOG files.
17:14:26.770	2015_10_14_17_13_58_packet_record.data contains data at freq : 7584.000Hz
17:14:26.772	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:14:28.783	Configure sinewaves generation: 2.0Vpp @7680.000 Hz
17:14:30.820	Time has been set to : 498154470
17:14:30.822	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154471
17:14:30.824	Trigger will be sent to Analog Discovery at : 498154473
17:14:44.865	We start to log RAW and LOG files.
17:15:10.867	We stop to log RAW and LOG files.
17:15:12.870	2015_10_14_17_14_44_packet_record.data contains data at freq : 7680.000Hz
17:15:12.872	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:15:14.885	Configure sinewaves generation: 2.0Vpp @7776.000 Hz
17:15:16.937	Time has been set to : 498154516
17:15:16.939	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154517
17:15:16.941	Trigger will be sent to Analog Discovery at : 498154519
17:15:30.994	We start to log RAW and LOG files.
17:15:57.0	We stop to log RAW and LOG files.
17:15:59.3	2015_10_14_17_15_30_packet_record.data contains data at freq : 7776.000Hz
17:15:59.5	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:16:01.14	Configure sinewaves generation: 2.0Vpp @7872.000 Hz
17:16:03.67	Time has been set to : 498154563
17:16:03.69	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154564
17:16:03.71	Trigger will be sent to Analog Discovery at : 498154566
17:16:17.112	We start to log RAW and LOG files.
17:16:43.114	We stop to log RAW and LOG files.
17:16:45.117	2015_10_14_17_16_17_packet_record.data contains data at freq : 7872.000Hz
17:16:45.120	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:16:47.138	Configure sinewaves generation: 2.0Vpp @7968.000 Hz
17:16:49.191	Time has been set to : 498154609
17:16:49.193	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154610
17:16:49.195	Trigger will be sent to Analog Discovery at : 498154612
17:17:03.236	We start to log RAW and LOG files.
17:17:29.238	We stop to log RAW and LOG files.
17:17:31.241	2015_10_14_17_17_03_packet_record.data contains data at freq : 7968.000Hz
17:17:31.243	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:17:33.260	Configure sinewaves generation: 2.0Vpp @8064.000 Hz
17:17:35.303	Time has been set to : 498154655
17:17:35.305	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154656
17:17:35.307	Trigger will be sent to Analog Discovery at : 498154658
17:17:49.359	We start to log RAW and LOG files.
17:18:15.362	We stop to log RAW and LOG files.
17:18:17.365	2015_10_14_17_17_49_packet_record.data contains data at freq : 8064.000Hz
17:18:17.367	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:18:19.385	Configure sinewaves generation: 2.0Vpp @8160.000 Hz
17:18:21.439	Time has been set to : 498154701
17:18:21.441	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154702
17:18:21.443	Trigger will be sent to Analog Discovery at : 498154704
17:18:35.484	We start to log RAW and LOG files.
17:19:01.486	We stop to log RAW and LOG files.
17:19:03.489	2015_10_14_17_18_35_packet_record.data contains data at freq : 8160.000Hz
17:19:03.491	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:19:05.510	Configure sinewaves generation: 2.0Vpp @8256.000 Hz
17:19:07.563	Time has been set to : 498154747
17:19:07.565	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154748
17:19:07.567	Trigger will be sent to Analog Discovery at : 498154750
17:19:21.608	We start to log RAW and LOG files.
17:19:47.610	We stop to log RAW and LOG files.
17:19:49.613	2015_10_14_17_19_21_packet_record.data contains data at freq : 8256.000Hz
17:19:49.615	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:19:51.627	Configure sinewaves generation: 2.0Vpp @8352.000 Hz
17:19:53.675	Time has been set to : 498154793
17:19:53.676	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154794
17:19:53.678	Trigger will be sent to Analog Discovery at : 498154796
17:20:07.719	We start to log RAW and LOG files.
17:20:33.721	We stop to log RAW and LOG files.
17:20:35.725	2015_10_14_17_20_07_packet_record.data contains data at freq : 8352.000Hz
17:20:35.727	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:20:37.740	Configure sinewaves generation: 2.0Vpp @8448.000 Hz
17:20:39.795	Time has been set to : 498154839
17:20:39.797	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154840
17:20:39.799	Trigger will be sent to Analog Discovery at : 498154842
17:20:53.840	We start to log RAW and LOG files.
17:21:19.842	We stop to log RAW and LOG files.
17:21:21.845	2015_10_14_17_20_53_packet_record.data contains data at freq : 8448.000Hz
17:21:21.847	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:21:23.865	Configure sinewaves generation: 2.0Vpp @8544.000 Hz
17:21:25.923	Time has been set to : 498154885
17:21:25.925	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154886
17:21:25.927	Trigger will be sent to Analog Discovery at : 498154888
17:21:39.967	We start to log RAW and LOG files.
17:22:05.972	We stop to log RAW and LOG files.
17:22:07.975	2015_10_14_17_21_39_packet_record.data contains data at freq : 8544.000Hz
17:22:07.977	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:22:09.987	Configure sinewaves generation: 2.0Vpp @8640.000 Hz
17:22:12.34	Time has been set to : 498154932
17:22:12.36	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154933
17:22:12.37	Trigger will be sent to Analog Discovery at : 498154935
17:22:26.90	We start to log RAW and LOG files.
17:22:52.92	We stop to log RAW and LOG files.
17:22:54.95	2015_10_14_17_22_26_packet_record.data contains data at freq : 8640.000Hz
17:22:54.97	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:22:56.114	Configure sinewaves generation: 2.0Vpp @8736.000 Hz
17:22:58.168	Time has been set to : 498154978
17:22:58.170	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498154979
17:22:58.172	Trigger will be sent to Analog Discovery at : 498154981
17:23:12.213	We start to log RAW and LOG files.
17:23:38.215	We stop to log RAW and LOG files.
17:23:40.218	2015_10_14_17_23_12_packet_record.data contains data at freq : 8736.000Hz
17:23:40.221	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:23:42.237	Configure sinewaves generation: 2.0Vpp @8832.000 Hz
17:23:44.291	Time has been set to : 498155024
17:23:44.293	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155025
17:23:44.295	Trigger will be sent to Analog Discovery at : 498155027
17:23:58.336	We start to log RAW and LOG files.
17:24:24.338	We stop to log RAW and LOG files.
17:24:26.341	2015_10_14_17_23_58_packet_record.data contains data at freq : 8832.000Hz
17:24:26.344	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:24:28.362	Configure sinewaves generation: 2.0Vpp @8928.000 Hz
17:24:30.406	Time has been set to : 498155070
17:24:30.408	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155071
17:24:30.410	Trigger will be sent to Analog Discovery at : 498155073
17:24:44.461	We start to log RAW and LOG files.
17:25:10.463	We stop to log RAW and LOG files.
17:25:12.467	2015_10_14_17_24_44_packet_record.data contains data at freq : 8928.000Hz
17:25:12.469	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:25:14.486	Configure sinewaves generation: 2.0Vpp @9024.000 Hz
17:25:16.529	Time has been set to : 498155116
17:25:16.531	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155117
17:25:16.533	Trigger will be sent to Analog Discovery at : 498155119
17:25:30.574	We start to log RAW and LOG files.
17:25:56.576	We stop to log RAW and LOG files.
17:25:58.580	2015_10_14_17_25_30_packet_record.data contains data at freq : 9024.000Hz
17:25:58.582	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:26:00.598	Configure sinewaves generation: 2.0Vpp @9120.000 Hz
17:26:02.642	Time has been set to : 498155162
17:26:02.644	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155163
17:26:02.645	Trigger will be sent to Analog Discovery at : 498155165
17:26:16.698	We start to log RAW and LOG files.
17:26:42.700	We stop to log RAW and LOG files.
17:26:44.703	2015_10_14_17_26_16_packet_record.data contains data at freq : 9120.000Hz
17:26:44.706	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:26:46.722	Configure sinewaves generation: 2.0Vpp @9216.000 Hz
17:26:48.777	Time has been set to : 498155208
17:26:48.779	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155209
17:26:48.781	Trigger will be sent to Analog Discovery at : 498155211
17:27:02.822	We start to log RAW and LOG files.
17:27:28.824	We stop to log RAW and LOG files.
17:27:30.827	2015_10_14_17_27_02_packet_record.data contains data at freq : 9216.000Hz
17:27:30.830	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:27:32.843	Configure sinewaves generation: 2.0Vpp @9312.000 Hz
17:27:34.889	Time has been set to : 498155254
17:27:34.891	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155255
17:27:34.893	Trigger will be sent to Analog Discovery at : 498155257
17:27:48.945	We start to log RAW and LOG files.
17:28:14.953	We stop to log RAW and LOG files.
17:28:16.957	2015_10_14_17_27_48_packet_record.data contains data at freq : 9312.000Hz
17:28:16.959	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:28:18.971	Configure sinewaves generation: 2.0Vpp @9408.000 Hz
17:28:21.11	Time has been set to : 498155301
17:28:21.13	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155302
17:28:21.15	Trigger will be sent to Analog Discovery at : 498155304
17:28:35.56	We start to log RAW and LOG files.
17:29:01.58	We stop to log RAW and LOG files.
17:29:03.61	2015_10_14_17_28_35_packet_record.data contains data at freq : 9408.000Hz
17:29:03.64	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:29:05.80	Configure sinewaves generation: 2.0Vpp @9504.000 Hz
17:29:07.124	Time has been set to : 498155347
17:29:07.126	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155348
17:29:07.128	Trigger will be sent to Analog Discovery at : 498155350
17:29:21.168	We start to log RAW and LOG files.
17:29:47.171	We stop to log RAW and LOG files.
17:29:49.174	2015_10_14_17_29_21_packet_record.data contains data at freq : 9504.000Hz
17:29:49.177	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:29:51.195	Configure sinewaves generation: 2.0Vpp @9600.000 Hz
17:29:53.238	Time has been set to : 498155393
17:29:53.239	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155394
17:29:53.241	Trigger will be sent to Analog Discovery at : 498155396
17:30:07.283	We start to log RAW and LOG files.
17:30:33.285	We stop to log RAW and LOG files.
17:30:35.289	2015_10_14_17_30_07_packet_record.data contains data at freq : 9600.000Hz
17:30:35.291	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:30:37.304	Configure sinewaves generation: 2.0Vpp @9696.000 Hz
17:30:39.348	Time has been set to : 498155439
17:30:39.350	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155440
17:30:39.352	Trigger will be sent to Analog Discovery at : 498155442
17:30:53.403	We start to log RAW and LOG files.
17:31:19.405	We stop to log RAW and LOG files.
17:31:21.408	2015_10_14_17_30_53_packet_record.data contains data at freq : 9696.000Hz
17:31:21.411	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:31:23.420	Configure sinewaves generation: 2.0Vpp @9792.000 Hz
17:31:25.463	Time has been set to : 498155485
17:31:25.465	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155486
17:31:25.467	Trigger will be sent to Analog Discovery at : 498155488
17:31:39.519	We start to log RAW and LOG files.
17:32:05.522	We stop to log RAW and LOG files.
17:32:07.525	2015_10_14_17_31_39_packet_record.data contains data at freq : 9792.000Hz
17:32:07.527	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:32:09.543	Configure sinewaves generation: 2.0Vpp @9888.000 Hz
17:32:11.596	Time has been set to : 498155531
17:32:11.598	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155532
17:32:11.600	Trigger will be sent to Analog Discovery at : 498155534
17:32:25.641	We start to log RAW and LOG files.
17:32:51.644	We stop to log RAW and LOG files.
17:32:53.647	2015_10_14_17_32_25_packet_record.data contains data at freq : 9888.000Hz
17:32:53.649	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:32:55.658	Configure sinewaves generation: 2.0Vpp @9984.000 Hz
17:32:57.709	Time has been set to : 498155577
17:32:57.711	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155578
17:32:57.713	Trigger will be sent to Analog Discovery at : 498155580
17:33:11.754	We start to log RAW and LOG files.
17:33:37.757	We stop to log RAW and LOG files.
17:33:39.760	2015_10_14_17_33_11_packet_record.data contains data at freq : 9984.000Hz
17:33:39.762	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:33:41.781	Configure sinewaves generation: 2.0Vpp @10080.000 Hz
17:33:43.825	Time has been set to : 498155623
17:33:43.827	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155624
17:33:43.829	Trigger will be sent to Analog Discovery at : 498155626
17:33:57.882	We start to log RAW and LOG files.
17:34:23.884	We stop to log RAW and LOG files.
17:34:25.888	2015_10_14_17_33_57_packet_record.data contains data at freq : 10080.000Hz
17:34:25.890	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:34:27.904	Configure sinewaves generation: 2.0Vpp @10176.000 Hz
17:34:29.945	Time has been set to : 498155669
17:34:29.946	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155670
17:34:29.948	Trigger will be sent to Analog Discovery at : 498155672
17:34:43.991	We start to log RAW and LOG files.
17:35:10.0	We stop to log RAW and LOG files.
17:35:12.3	2015_10_14_17_34_43_packet_record.data contains data at freq : 10176.000Hz
17:35:12.6	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:35:14.14	Configure sinewaves generation: 2.0Vpp @10272.000 Hz
17:35:16.67	Time has been set to : 498155716
17:35:16.69	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155717
17:35:16.71	Trigger will be sent to Analog Discovery at : 498155719
17:35:30.111	We start to log RAW and LOG files.
17:35:56.114	We stop to log RAW and LOG files.
17:35:58.117	2015_10_14_17_35_30_packet_record.data contains data at freq : 10272.000Hz
17:35:58.120	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:36:00.127	Configure sinewaves generation: 2.0Vpp @10368.000 Hz
17:36:02.169	Time has been set to : 498155762
17:36:02.171	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155763
17:36:02.173	Trigger will be sent to Analog Discovery at : 498155765
17:36:16.214	We start to log RAW and LOG files.
17:36:42.216	We stop to log RAW and LOG files.
17:36:44.220	2015_10_14_17_36_16_packet_record.data contains data at freq : 10368.000Hz
17:36:44.222	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:36:46.237	Configure sinewaves generation: 2.0Vpp @10464.000 Hz
17:36:48.280	Time has been set to : 498155808
17:36:48.282	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155809
17:36:48.285	Trigger will be sent to Analog Discovery at : 498155811
17:37:02.336	We start to log RAW and LOG files.
17:37:28.339	We stop to log RAW and LOG files.
17:37:30.342	2015_10_14_17_37_02_packet_record.data contains data at freq : 10464.000Hz
17:37:30.345	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:37:32.361	Configure sinewaves generation: 2.0Vpp @10560.000 Hz
17:37:34.406	Time has been set to : 498155854
17:37:34.408	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155855
17:37:34.410	Trigger will be sent to Analog Discovery at : 498155857
17:37:48.451	We start to log RAW and LOG files.
17:38:14.453	We stop to log RAW and LOG files.
17:38:16.457	2015_10_14_17_37_48_packet_record.data contains data at freq : 10560.000Hz
17:38:16.459	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:38:18.475	Configure sinewaves generation: 2.0Vpp @10656.000 Hz
17:38:20.530	Time has been set to : 498155900
17:38:20.532	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155901
17:38:20.534	Trigger will be sent to Analog Discovery at : 498155903
17:38:34.575	We start to log RAW and LOG files.
17:39:00.577	We stop to log RAW and LOG files.
17:39:02.580	2015_10_14_17_38_34_packet_record.data contains data at freq : 10656.000Hz
17:39:02.583	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:39:04.594	Configure sinewaves generation: 2.0Vpp @10752.000 Hz
17:39:06.643	Time has been set to : 498155946
17:39:06.645	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155947
17:39:06.647	Trigger will be sent to Analog Discovery at : 498155949
17:39:20.699	We start to log RAW and LOG files.
17:39:46.701	We stop to log RAW and LOG files.
17:39:48.705	2015_10_14_17_39_20_packet_record.data contains data at freq : 10752.000Hz
17:39:48.707	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:39:50.716	Configure sinewaves generation: 2.0Vpp @10848.000 Hz
17:39:52.765	Time has been set to : 498155992
17:39:52.767	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498155993
17:39:52.769	Trigger will be sent to Analog Discovery at : 498155995
17:40:06.821	We start to log RAW and LOG files.
17:40:32.824	We stop to log RAW and LOG files.
17:40:34.827	2015_10_14_17_40_06_packet_record.data contains data at freq : 10848.000Hz
17:40:34.830	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:40:36.839	Configure sinewaves generation: 2.0Vpp @10944.000 Hz
17:40:38.877	Time has been set to : 498156038
17:40:38.879	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156039
17:40:38.881	Trigger will be sent to Analog Discovery at : 498156041
17:40:52.923	We start to log RAW and LOG files.
17:41:18.926	We stop to log RAW and LOG files.
17:41:20.930	2015_10_14_17_40_52_packet_record.data contains data at freq : 10944.000Hz
17:41:20.932	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:41:22.940	Configure sinewaves generation: 2.0Vpp @11040.000 Hz
17:41:24.978	Time has been set to : 498156084
17:41:24.981	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156085
17:41:24.983	Trigger will be sent to Analog Discovery at : 498156087
17:41:39.33	We start to log RAW and LOG files.
17:42:05.44	We stop to log RAW and LOG files.
17:42:07.47	2015_10_14_17_41_39_packet_record.data contains data at freq : 11040.000Hz
17:42:07.50	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:42:09.58	Configure sinewaves generation: 2.0Vpp @11136.000 Hz
17:42:11.112	Time has been set to : 498156131
17:42:11.114	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156132
17:42:11.116	Trigger will be sent to Analog Discovery at : 498156134
17:42:25.157	We start to log RAW and LOG files.
17:42:51.159	We stop to log RAW and LOG files.
17:42:53.163	2015_10_14_17_42_25_packet_record.data contains data at freq : 11136.000Hz
17:42:53.165	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:42:55.182	Configure sinewaves generation: 2.0Vpp @11232.000 Hz
17:42:57.224	Time has been set to : 498156177
17:42:57.226	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156178
17:42:57.228	Trigger will be sent to Analog Discovery at : 498156180
17:43:11.280	We start to log RAW and LOG files.
17:43:37.283	We stop to log RAW and LOG files.
17:43:39.286	2015_10_14_17_43_11_packet_record.data contains data at freq : 11232.000Hz
17:43:39.288	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:43:41.305	Configure sinewaves generation: 2.0Vpp @11328.000 Hz
17:43:43.347	Time has been set to : 498156223
17:43:43.349	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156224
17:43:43.351	Trigger will be sent to Analog Discovery at : 498156226
17:43:57.404	We start to log RAW and LOG files.
17:44:23.406	We stop to log RAW and LOG files.
17:44:25.410	2015_10_14_17_43_57_packet_record.data contains data at freq : 11328.000Hz
17:44:25.412	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:44:27.420	Configure sinewaves generation: 2.0Vpp @11424.000 Hz
17:44:29.474	Time has been set to : 498156269
17:44:29.476	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156270
17:44:29.478	Trigger will be sent to Analog Discovery at : 498156272
17:44:43.519	We start to log RAW and LOG files.
17:45:09.522	We stop to log RAW and LOG files.
17:45:11.525	2015_10_14_17_44_43_packet_record.data contains data at freq : 11424.000Hz
17:45:11.527	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:45:13.544	Configure sinewaves generation: 2.0Vpp @11520.000 Hz
17:45:15.587	Time has been set to : 498156315
17:45:15.589	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156316
17:45:15.591	Trigger will be sent to Analog Discovery at : 498156318
17:45:29.632	We start to log RAW and LOG files.
17:45:55.634	We stop to log RAW and LOG files.
17:45:57.638	2015_10_14_17_45_29_packet_record.data contains data at freq : 11520.000Hz
17:45:57.640	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:45:59.651	Configure sinewaves generation: 2.0Vpp @11616.000 Hz
17:46:01.699	Time has been set to : 498156361
17:46:01.701	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156362
17:46:01.703	Trigger will be sent to Analog Discovery at : 498156364
17:46:15.755	We start to log RAW and LOG files.
17:46:41.758	We stop to log RAW and LOG files.
17:46:43.761	2015_10_14_17_46_15_packet_record.data contains data at freq : 11616.000Hz
17:46:43.764	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:46:45.776	Configure sinewaves generation: 2.0Vpp @11712.000 Hz
17:46:47.833	Time has been set to : 498156407
17:46:47.835	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156408
17:46:47.837	Trigger will be sent to Analog Discovery at : 498156410
17:47:01.878	We start to log RAW and LOG files.
17:47:27.880	We stop to log RAW and LOG files.
17:47:29.884	2015_10_14_17_47_01_packet_record.data contains data at freq : 11712.000Hz
17:47:29.886	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:47:31.898	Configure sinewaves generation: 2.0Vpp @11808.000 Hz
17:47:33.945	Time has been set to : 498156453
17:47:33.947	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156454
17:47:33.949	Trigger will be sent to Analog Discovery at : 498156456
17:47:47.990	We start to log RAW and LOG files.
17:48:13.997	We stop to log RAW and LOG files.
17:48:16.1	2015_10_14_17_47_47_packet_record.data contains data at freq : 11808.000Hz
17:48:16.3	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:48:18.14	Configure sinewaves generation: 2.0Vpp @11904.000 Hz
17:48:20.67	Time has been set to : 498156500
17:48:20.69	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156501
17:48:20.71	Trigger will be sent to Analog Discovery at : 498156503
17:48:34.123	We start to log RAW and LOG files.
17:49:00.125	We stop to log RAW and LOG files.
17:49:02.129	2015_10_14_17_48_34_packet_record.data contains data at freq : 11904.000Hz
17:49:02.131	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:49:04.149	Configure sinewaves generation: 2.0Vpp @12000.000 Hz
17:49:06.200	Time has been set to : 498156546
17:49:06.202	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156547
17:49:06.204	Trigger will be sent to Analog Discovery at : 498156549
17:49:20.245	We start to log RAW and LOG files.
17:49:46.247	We stop to log RAW and LOG files.
17:49:48.251	2015_10_14_17_49_20_packet_record.data contains data at freq : 12000.000Hz
17:49:48.253	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:49:50.271	Configure sinewaves generation: 2.0Vpp @12096.000 Hz
17:49:52.327	Time has been set to : 498156592
17:49:52.329	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156593
17:49:52.331	Trigger will be sent to Analog Discovery at : 498156595
17:50:06.374	We start to log RAW and LOG files.
17:50:32.383	We stop to log RAW and LOG files.
17:50:34.387	2015_10_14_17_50_06_packet_record.data contains data at freq : 12096.000Hz
17:50:34.389	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:50:36.399	Configure sinewaves generation: 2.0Vpp @12192.000 Hz
17:50:38.441	Time has been set to : 498156638
17:50:38.443	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156639
17:50:38.445	Trigger will be sent to Analog Discovery at : 498156641
17:50:52.496	We start to log RAW and LOG files.
17:51:18.498	We stop to log RAW and LOG files.
17:51:20.502	2015_10_14_17_50_52_packet_record.data contains data at freq : 12192.000Hz
17:51:20.504	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:51:22.522	Configure sinewaves generation: 2.0Vpp @12288.000 Hz
17:51:24.576	Time has been set to : 498156684
17:51:24.578	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156685
17:51:24.580	Trigger will be sent to Analog Discovery at : 498156687
17:51:38.621	We start to log RAW and LOG files.
17:52:04.623	We stop to log RAW and LOG files.
17:52:06.627	2015_10_14_17_51_38_packet_record.data contains data at freq : 12288.000Hz
17:52:06.629	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:52:08.644	Configure sinewaves generation: 2.0Vpp @13056.000 Hz
17:52:10.699	Time has been set to : 498156730
17:52:10.701	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156731
17:52:10.703	Trigger will be sent to Analog Discovery at : 498156733
17:52:24.744	We start to log RAW and LOG files.
17:52:50.746	We stop to log RAW and LOG files.
17:52:52.750	2015_10_14_17_52_24_packet_record.data contains data at freq : 13056.000Hz
17:52:52.752	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:52:54.768	Configure sinewaves generation: 2.0Vpp @13824.000 Hz
17:52:56.812	Time has been set to : 498156776
17:52:56.814	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156777
17:52:56.816	Trigger will be sent to Analog Discovery at : 498156779
17:53:10.868	We start to log RAW and LOG files.
17:53:36.870	We stop to log RAW and LOG files.
17:53:38.874	2015_10_14_17_53_10_packet_record.data contains data at freq : 13824.000Hz
17:53:38.876	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:53:40.892	Configure sinewaves generation: 2.0Vpp @14592.000 Hz
17:53:42.934	Time has been set to : 498156822
17:53:42.937	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156823
17:53:42.939	Trigger will be sent to Analog Discovery at : 498156825
17:53:56.991	We start to log RAW and LOG files.
17:54:22.994	We stop to log RAW and LOG files.
17:54:24.998	2015_10_14_17_53_56_packet_record.data contains data at freq : 14592.000Hz
17:54:25.0	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:54:27.13	Configure sinewaves generation: 2.0Vpp @15360.000 Hz
17:54:29.68	Time has been set to : 498156869
17:54:29.70	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156870
17:54:29.72	Trigger will be sent to Analog Discovery at : 498156872
17:54:43.113	We start to log RAW and LOG files.
17:55:09.116	We stop to log RAW and LOG files.
17:55:11.120	2015_10_14_17_54_43_packet_record.data contains data at freq : 15360.000Hz
17:55:11.122	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:55:13.137	Configure sinewaves generation: 2.0Vpp @16128.000 Hz
17:55:15.180	Time has been set to : 498156915
17:55:15.182	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156916
17:55:15.184	Trigger will be sent to Analog Discovery at : 498156918
17:55:29.224	We start to log RAW and LOG files.
17:55:55.227	We stop to log RAW and LOG files.
17:55:57.231	2015_10_14_17_55_29_packet_record.data contains data at freq : 16128.000Hz
17:55:57.233	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:55:59.250	Configure sinewaves generation: 2.0Vpp @16896.000 Hz
17:56:01.303	Time has been set to : 498156961
17:56:01.306	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498156962
17:56:01.308	Trigger will be sent to Analog Discovery at : 498156964
17:56:15.348	We start to log RAW and LOG files.
17:56:41.351	We stop to log RAW and LOG files.
17:56:43.355	2015_10_14_17_56_15_packet_record.data contains data at freq : 16896.000Hz
17:56:43.357	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:56:45.366	Configure sinewaves generation: 2.0Vpp @17664.000 Hz
17:56:47.408	Time has been set to : 498157007
17:56:47.411	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157008
17:56:47.413	Trigger will be sent to Analog Discovery at : 498157010
17:57:01.465	We start to log RAW and LOG files.
17:57:27.467	We stop to log RAW and LOG files.
17:57:29.471	2015_10_14_17_57_01_packet_record.data contains data at freq : 17664.000Hz
17:57:29.473	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:57:31.487	Configure sinewaves generation: 2.0Vpp @18432.000 Hz
17:57:33.533	Time has been set to : 498157053
17:57:33.535	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157054
17:57:33.537	Trigger will be sent to Analog Discovery at : 498157056
17:57:47.578	We start to log RAW and LOG files.
17:58:13.581	We stop to log RAW and LOG files.
17:58:15.585	2015_10_14_17_57_47_packet_record.data contains data at freq : 18432.000Hz
17:58:15.587	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:58:17.602	Configure sinewaves generation: 2.0Vpp @19200.000 Hz
17:58:19.654	Time has been set to : 498157099
17:58:19.656	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157100
17:58:19.658	Trigger will be sent to Analog Discovery at : 498157102
17:58:33.699	We start to log RAW and LOG files.
17:58:59.701	We stop to log RAW and LOG files.
17:59:01.705	2015_10_14_17_58_33_packet_record.data contains data at freq : 19200.000Hz
17:59:01.707	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:59:03.716	Configure sinewaves generation: 2.0Vpp @19968.000 Hz
17:59:05.759	Time has been set to : 498157145
17:59:05.762	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157146
17:59:05.764	Trigger will be sent to Analog Discovery at : 498157148
17:59:19.815	We start to log RAW and LOG files.
17:59:45.817	We stop to log RAW and LOG files.
17:59:47.821	2015_10_14_17_59_19_packet_record.data contains data at freq : 19968.000Hz
17:59:47.824	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
17:59:49.841	Configure sinewaves generation: 2.0Vpp @20736.000 Hz
17:59:51.884	Time has been set to : 498157191
17:59:51.886	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157192
17:59:51.888	Trigger will be sent to Analog Discovery at : 498157194
18:00:05.941	We start to log RAW and LOG files.
18:00:31.945	We stop to log RAW and LOG files.
18:00:33.949	2015_10_14_18_00_05_packet_record.data contains data at freq : 20736.000Hz
18:00:33.952	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:00:35.962	Configure sinewaves generation: 2.0Vpp @21504.000 Hz
18:00:38.11	Time has been set to : 498157238
18:00:38.13	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157239
18:00:38.15	Trigger will be sent to Analog Discovery at : 498157241
18:00:52.56	We start to log RAW and LOG files.
18:01:18.63	We stop to log RAW and LOG files.
18:01:20.67	2015_10_14_18_00_52_packet_record.data contains data at freq : 21504.000Hz
18:01:20.70	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:01:22.81	Configure sinewaves generation: 2.0Vpp @22272.000 Hz
18:01:24.112	Time has been set to : 498157284
18:01:24.115	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157285
18:01:24.117	Trigger will be sent to Analog Discovery at : 498157287
18:01:38.168	We start to log RAW and LOG files.
18:02:04.171	We stop to log RAW and LOG files.
18:02:06.174	2015_10_14_18_01_38_packet_record.data contains data at freq : 22272.000Hz
18:02:06.177	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:02:08.194	Configure sinewaves generation: 2.0Vpp @23040.000 Hz
18:02:10.247	Time has been set to : 498157330
18:02:10.249	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157331
18:02:10.251	Trigger will be sent to Analog Discovery at : 498157333
18:02:24.292	We start to log RAW and LOG files.
18:02:50.294	We stop to log RAW and LOG files.
18:02:52.298	2015_10_14_18_02_24_packet_record.data contains data at freq : 23040.000Hz
18:02:52.300	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:02:54.318	Configure sinewaves generation: 2.0Vpp @23808.000 Hz
18:02:56.359	Time has been set to : 498157376
18:02:56.361	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157377
18:02:56.363	Trigger will be sent to Analog Discovery at : 498157379
18:03:10.405	We start to log RAW and LOG files.
18:03:36.408	We stop to log RAW and LOG files.
18:03:38.412	2015_10_14_18_03_10_packet_record.data contains data at freq : 23808.000Hz
18:03:38.414	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:03:40.422	Configure sinewaves generation: 2.0Vpp @24576.000 Hz
18:03:42.475	Time has been set to : 498157422
18:03:42.477	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157423
18:03:42.480	Trigger will be sent to Analog Discovery at : 498157425
18:03:56.520	We start to log RAW and LOG files.
18:04:22.523	We stop to log RAW and LOG files.
18:04:24.527	2015_10_14_18_03_56_packet_record.data contains data at freq : 24576.000Hz
18:04:24.529	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:04:26.545	Configure sinewaves generation: 2.0Vpp @25344.000 Hz
18:04:28.589	Time has been set to : 498157468
18:04:28.591	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157469
18:04:28.593	Trigger will be sent to Analog Discovery at : 498157471
18:04:42.644	We start to log RAW and LOG files.
18:05:08.647	We stop to log RAW and LOG files.
18:05:10.651	2015_10_14_18_04_42_packet_record.data contains data at freq : 25344.000Hz
18:05:10.653	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:05:12.661	Configure sinewaves generation: 2.0Vpp @26112.000 Hz
18:05:14.711	Time has been set to : 498157514
18:05:14.713	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157515
18:05:14.715	Trigger will be sent to Analog Discovery at : 498157517
18:05:28.756	We start to log RAW and LOG files.
18:05:54.758	We stop to log RAW and LOG files.
18:05:56.762	2015_10_14_18_05_28_packet_record.data contains data at freq : 26112.000Hz
18:05:56.765	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:05:58.773	Configure sinewaves generation: 2.0Vpp @26880.000 Hz
18:06:00.824	Time has been set to : 498157560
18:06:00.826	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157561
18:06:00.828	Trigger will be sent to Analog Discovery at : 498157563
18:06:14.880	We start to log RAW and LOG files.
18:06:40.883	We stop to log RAW and LOG files.
18:06:42.887	2015_10_14_18_06_14_packet_record.data contains data at freq : 26880.000Hz
18:06:42.889	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:06:44.898	Configure sinewaves generation: 2.0Vpp @27648.000 Hz
18:06:46.948	Time has been set to : 498157606
18:06:46.950	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157607
18:06:46.952	Trigger will be sent to Analog Discovery at : 498157609
18:07:00.993	We start to log RAW and LOG files.
18:07:26.997	We stop to log RAW and LOG files.
18:07:29.0	2015_10_14_18_07_00_packet_record.data contains data at freq : 27648.000Hz
18:07:29.3	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:07:31.12	Configure sinewaves generation: 2.0Vpp @28416.000 Hz
18:07:33.56	Time has been set to : 498157653
18:07:33.58	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157654
18:07:33.60	Trigger will be sent to Analog Discovery at : 498157656
18:07:47.112	We start to log RAW and LOG files.
18:08:13.115	We stop to log RAW and LOG files.
18:08:15.118	2015_10_14_18_07_47_packet_record.data contains data at freq : 28416.000Hz
18:08:15.121	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:08:17.137	Configure sinewaves generation: 2.0Vpp @29184.000 Hz
18:08:19.180	Time has been set to : 498157699
18:08:19.182	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157700
18:08:19.184	Trigger will be sent to Analog Discovery at : 498157702
18:08:33.236	We start to log RAW and LOG files.
18:08:59.239	We stop to log RAW and LOG files.
18:09:01.243	2015_10_14_18_08_33_packet_record.data contains data at freq : 29184.000Hz
18:09:01.245	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:09:03.260	Configure sinewaves generation: 2.0Vpp @29952.000 Hz
18:09:05.303	Time has been set to : 498157745
18:09:05.305	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157746
18:09:05.307	Trigger will be sent to Analog Discovery at : 498157748
18:09:19.348	We start to log RAW and LOG files.
18:09:45.350	We stop to log RAW and LOG files.
18:09:47.354	2015_10_14_18_09_19_packet_record.data contains data at freq : 29952.000Hz
18:09:47.356	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:09:49.370	Configure sinewaves generation: 2.0Vpp @30720.000 Hz
18:09:51.420	Time has been set to : 498157791
18:09:51.422	TC_LFR_ENTER_MODE *** put LFR in NORMAL mode AT 498157792
18:09:51.424	Trigger will be sent to Analog Discovery at : 498157794
18:10:05.476	We start to log RAW and LOG files.
18:10:31.479	We stop to log RAW and LOG files.
18:10:33.483	2015_10_14_18_10_05_packet_record.data contains data at freq : 30720.000Hz
18:10:33.485	TC_LFR_ENTER_MODE *** put LFR in STANDBY mode
18:10:35.501	end of the test